## Slide 1: Applications, Developments, and Recent Work with O- Beams Noel S Smith and Stephane Melanson Oregon Physics LLC
n.smith@oregon-physics.com
OREGON PHYSICS LLC, 15236 NW Greenbrier Parkway, Beaverton, Oregon 97006, USA

## Slide 2: Overview
What is Secondary Ion Mass Spectrometry (SIMS)
Why Focused O- beams are used.
Ion Source requirements for Focused Ion Beams
Evolution from Duoplasmatrons to ICP ion sources.
Possible Future Developments

## Slide 3: Secondary Ion Mass Spectrometry (SIMS)
Hit a surface with a focused keV ion beam, collect the few sputtered atoms that leave as ions, and mass-analyse them — the most sensitive surface analysis technique there is.
Sample being analyzed
primary
objective lens
ion
source
primary beam:
O⁻, O₂⁺ or Cs⁺, 1–30 keV
focused and rastered
secondary immersion /
extraction lens
sample at ±kV
secondary ions M⁺ (or M⁻)
10⁻⁴–10⁻¹ of what is sputtered;
accelerated to keV by the sample bias
magnetic sector
(mass spectrometer)
detector
counts ions
one m/z at a time reaches the detector
Why use SIMS?
▪ ppm–ppb sensitivity; all elements incl. H, Li, B
▪ isotopes resolved → dating, tracers, meteorites
▪ depth profiles with ~1 nm resolution (dopants)
▪ imaging at 50 nm – 10 µm (NanoSIMS, ToF-SIMS)
However!
▪ Only a tiny, chemistry-dependent fraction is ionised → primary ion: O for M⁺, Cs for M⁻
▪ resolution, speed, depth resolution follow from spot size & current → source brightness
Output:
Mass spectrum
Depth profile
Element map
Instrument families: Magnetic Sector SIMS (Cameca NanoSIMS/1300/IMS Series — magnetic sector, high mass resolution and transmission – nuclear, biology, geology, etc..)
ToF-SIMS (pulsed Bi⁺, Ar cluster — molecular surface info)
Quadrupole SIMS (mostly semiconductor applications)

## Slide 4: Why Oxygen Primary Beams for SIMS?
sample surface, oxygen-saturated
O
O
O
O
O
M
M
M
M
O⁻ / O2+ primary
(implants O into the surface)
M+ secondary ion
(M–O bond breaks, e⁻ stays on O)
“Bond-breaking” picture of the oxygen matrix effect
Secondary ion yield of many electropositive species enhanced by implanted oxygen
Sputtering removes atoms efficiently; the problem is that almost all leave neutral
Oxygen in the surface raises the work function and forms M–O bonds; sputtered matrix atoms leave as M+ with 10–1000× higher probability
The beam itself implants oxygen, so the surface reaches a steady oxide state.
Complement: Cs+ lowers the work function → boosts negative secondaries (H, C, O, S, Si, halogens, As, P)
Oxygen beams ⇒ positive secondaries ⇒ the electropositive elements
Dopants and metals in semiconductors: B, Al, W, Cu, Fe, Ti
Geoscience, cosmochemistry, nuclear forensics: Mg, Al, Si, Ca, Ti, Cr, Fe, rare earths, U, Th, Pb
Biology: Ca, Mg, Na, K and Lanthanides

## Slide 5: Secondary ion yield enhancement of electropositive ions with oxygenated surface
92
U
93
Np
94
Pu
95
Am
96
Cm
U–Cm shaded here (unshaded in the original only because no RSF data existed): the actinides have low ionisation potentials (U 6.2 eV, Pu 6.0 eV) and form very polar oxides, so O⁻ / O₂⁺ is the standard beam for U and Pu isotopics — measured as U⁺, UO⁺ and UO₂⁺.
Lanthanides
Actinides
Alkali metals and alikali earth metals give highest ion yield

## Slide 6: SIMS for Trace Element Detection
Primary ion beam
O⁻, O2+, Cs+
1–30 keV, pA–µA
50 nm–µm spot
Sputter
from top 1–2 nm of
the sample
Secondary ions
~0.01%–10%
of sputtered atoms
leave charged
Mass spectrometer
magnetic sector,
ToF or quadrupole
Result
isotope / element
map or depth profile

## Slide 7: Why O⁻ for Insulating Samples?
Positive primary (O₂⁺)
extraction electrode 0 V
insulating sample on a +5 kV holder
O₂⁺
secondary e⁻ (few eV)
pulled straight back in
+ + + + + + +
→ surface potential rises without limit
▪ primary beam is deflected and defocused → image degrades
▪ secondary ions gain energy → shifted out of the spectrometer’s energy window
Negative primary (O⁻)
extraction electrode 0 V
insulating sample on a +5 kV holder
O⁻
brings −
M⁺ out
secondary e⁻ (few eV)
land nearby
−
e⁻ e⁻ e⁻ e⁻ e⁻
net charge per ion: −1, balanced by ≈ 1 escaping electron
→ spot settles a few volts negative and stays there
▪ O⁻ in and M⁺ out drive the surface slightly negative
which ejects secondary e⁻ into the surrounding area.
Secondary e- are only a few eV, so they stay close and re-land, but analysis region becomes stable.
A negative primary brings its own neutralisation.
No low energy electron flood gun needed.
Stable images and spectra.

## Slide 8: O- Primary Ions Can Share an Objective lens with Positive Secondary Ions
Conventional SIMS: two separate objective lenses
sample
Secondary ion beam extraction lens
Primary ion beam
objective lens
lens bodies nearly touch:
same working distance, but the
primary must arrive obliquely
Lens working distance (WD)≈ 5-10 mm
sample at a few kV,
oblique primary impact
Coaxial (NanoSIMS): one objective for both beams
sample
immersion objective: primary in, secondaries out
primary ↓
↑ secondary
gap ≈ 0.4 mm
(both beams)
normal incidence.
sample and primary of opposite polarity (increased impact energy)
Actual Cameca NanoSIMS lens is more complicated and provides some independence of primary and secondary focusing.
High secondary ion collection efficiency
Primary probe size
aberrations grow with working distance:
Cs, Cc scale with WD → dmin ↑
WD = 5–10 mm vs 0.4 mm
Extraction field at the sample
E = V / gap: 5 kV / 5 mm ≈ 1 V/µm
5kV / 0.4 mm ≈ 12.5 V/µm
stronger field → increased secondary ion collection efficiency
Beam Polarity
The two beams must have opposite sign,
O⁻ required for positive secondaries
Approximately:
Cc ∝ f
Cs ∝ f³

## Slide 9: How a Primary Ion Column Works
Ion source
virtual source dv
angular intensity I′
energy E, spread ΔE
Br = I′ / (π dv² E)
Primary ion optics
demagnification M · aberrations Cs, Cc
for M ≪ 1 only the objective lens matters
Focused spot
size d
current I
half-angle α
Three blurs, added in quadrature
α = √( I / (π · I′ · M²) )
dg = M · dv
dc = (ΔE / E) · Cc · α
ds = ½ · Cs · α³
d = √( dg² + dc² + ds² )
Half-Angle needed for current, I
Gaussian image
Chromatic blur
Spherical blur
Total spot
The source fixes Br and ΔE — nothing downstream can improve them
The column fixes Cs, Cc (design) and M (lens excitation — the one knob you turn)
Ask for more current → larger α → aberrations grow; demagnify harder → smaller image but larger α again
So for every current there is an optimum M
"Classical (quadrature) estimate; Barth & Kruit, Optik 101 (1996) give a more accurate FW50 addition rule."

## Slide 10: Minimum Spot: Balancing the Three Blurs
Chart: Magnification, (M)
| 1 | 0.66666666666666663 | 0.36363636363636376 | 0.21052631578947376 | 0.14814814814814822 | 0.11428571428571435 | 9.3023255813953529E-2 | 6.7796610169491581E-2 | 5.3333333333333385E-2 | 4.8192771084337387E-2
 | 50 | 33.3333 | 18.1818 | 10.5263 | 7.4074 | 5.7143 | 4.6512 | 3.3898 | 2.6667 | 2.4096
 | 0.077324 | 0.2358 | 1.3482 | 6.2779 | 16.8447 | 35.0605 | 62.9373 | 155.7214 | 311.2921 | 417.6523
 | 6.4284 | 7.0384 | 10.2554 | 16.6259 | 23.3606 | 30.2141 | 37.1211 | 51.0094 | 64.9444 | 71.921
 | 50.4116 | 34.0691 | 20.9182 | 20.6552 | 29.7377 | 46.6346 | 73.2169 | 163.8981 | 318.0058 | 423.8064
Sweep M at a fixed current
Gaussian image ∝ M, chromatic ∝ 1/M, spherical ∝ 1/M³
The quadrature sum has a minimum → that M is the optimum lens excitation
Repeat for each current → the spot-size-vs-current curve shown later for Hyperion II
What the source buys you
At low I (chromatic-limited):
dmin ∝ (Cc ΔE/E)½ · (I / Br)¼
dmin scales with the inverse 4th root of the brightness, such that 10x gain in brightness results a factor of 1.8 reduction on spot size.
Reducing the energy spread by a factor of 2 and you get a factor of 1.4 reduction in dmin.

## Slide 11: ~1965-2015 - Only Duoplasmatrons for O2+ and O- Beams
e
e
e
e
e
e
e
e
e
e
e
+
+
+
+
+
+
+
~ -30V
~ -500V
Long, Coathe and Drummond – Cambridge Univ – 1995
Duoplasmatron
Highest brightness for O2+ , O- and inert gases.
For high brightness, you suffer with the short and unpredictable lifetimes.
Duoplasmatron | 
Brightness (Am-2sr-1V-1) | ~500 (O2+) and ~5 (O-)
Energy Spread (eV) | 10-15eV
Lifetime | 1 week – 6 months
Beam Stability | ±1% over 30 minutes
Gases | oxygen, hydrogen, inert

## Slide 12: Cold Plasma Creation with Inductive Power Coupling
Time varying current in an external coil creates a time varying B field, which induces an azimuthal E field.
xE=-B/t
Induction E field accelerates electrons to ionize resident gas.
(when Xe plasma density, no=1x1012cm-3)
~9GHz
~18MHz
10

## Slide 13: Ion Sources for Focused Ion Beams
>1x1012cm-3








Virtual Source
dv~5-50m
+10kV
0kV
Radius (um)
Angle (mrads)
Plasma
Source Electrode (250um aperture)
Extractor Electrode

## Slide 14
O- for SIMS imaging of geological, semiconductor and biological materials
PLASMA
Plasma Chamber
with low loss tangent
Faraday Shield
High Azimuthal Impedance
Extractor Electrode
Pole Gap
~600 Guass
Magnetic Poles
Transverse B
RF Antenna
Ion-Ion plasma
(depleted of hot electrons)
Extracting Negative Ions (High Res. SIMS Imaging)

## Slide 15
Oxygen Plasma
O- formed primarily by dissociative electron attachment.
O2 + e− → O + O−
O- Ion Source Parameters
 | Duoplasmatron | ICP (Hyperion)
Energy Normalized Brightness (Am-2sr-1V-1) | 5 | 100
Axial Energy Spread (eV) | 10-15 | 3.5
Lifetime | 1 week – 6 months | >2 years

## Slide 16: Hyperion II
Hyperion II - Parameters
  0 to ±30kV
  Typically operated with Xe, O2, He or N2
  Energy Normalized Brightness
    + ions Xe: 1x104, O2 : 4x103, He: 6.5x103 Am-2sr-1V-1
    - ions O :1x102 Am-2sr-1V-1
  Virtual source (dv)
    5um or 12um (+ ions)
    50-70um (- ions)
  Axial Energy Spread
    5eV (+ ions)
    3.5eV (- ions)
  Lifetime > 2 years
  Stability <1% drift/ hour, <5% drift/24 hours

## Slide 17: SIMS Imaging Performance at LLNL
  High Imaging Resolution Trace Element SIMS
  Direct replacement for duoplasmatron on Cameca NanoSIMS.
  Image resolution, with O- reduced from 200nm to 50nm.
  Lifetime increased from 1 week to >2 years.
Hyperion II + LLNL NanoSIMS
Supported by DoE and Glenn Seaborg Inst.
(P. Weber, M. Kristo and I.Hutcheon)

## Slide 18: Hyperon II on Cameca NanoSIMS50
Data Courtesy of:
Cameca (F. Horreard)
Lawrence Livermore (P. Weber)
U.West. Aus. (M. Kilburn)
Max Planck (P. Hoppe)
~200pA
6pA

## Slide 19: Beam Current Stability
No sputtering in plasma source
With steady gas flow and temperature regulation, beam current is very stable.
Residual drift is primarily due to lab. temperature drift (~0.5%/ 0C)
11

## Slide 20: Possible Useful Developments
High Brightness
  Higher throughput
  Higher spatial resolution imaging
Molecular oxygen ions (O2- and O3-)
  Higher Depth Resolution
  Increased Useful Ion Yield
O3- provides ~2x gain in useful ion yield for U isotope analysis*
* ‘Improved Uranium Particle Analysis by SIMS using O3- Primary Ions’ E. Groopman et al,, NIST, J. Anal. At. Spectrom., 2022, 37(10)
 | O- | O2- | O3-
Typical | ~80% | ~19% | <1%
NIST | 67% | 20% | 13%

## Slide 21
THANKYOU!