## Slide 1: How Low Can We Go? Routine Reliability and Low Energy Limits for AMS with a Cs Sputter Source
Lauren T Bezzina
C. Vockenhuber, P. Gautschi, L. Wacker, U. Ramsperger, A. Müller
Laboratory of Ion Beam Physics, ETH Zürich
NIBS2026 – September 2026

## Slide 2
AMS – Why do we do it?
  AMS is an ultra-sensitive method to detect trace amounts of long-lived radionuclides.
  With both cosmogenic and anthropogenic sources, these nuclides can used as a ‘clock’ for:
  Archeology/paleoclimate (14C 5.7 x 103 yr)
  Earth sciences/geomorphology – 10Be (1.38 x 106 yr), 26Al (717 x 103 yr)
  Environmental/hydrology — 36Cl (300 x 103 yr), 129I (15.7 x106 yr)
  Biomedicine — 14C, 41Ca (99 x 103 yr)
  Nuclear science/oceanography — actinides (Pu isotopes, U, etc.)
1

## Slide 3
AMS – Why do we do it?
  AMS is an ultra-sensitive method to detect trace amounts of long-lived radionuclides.
  With both cosmogenic and anthropogenic sources, these nuclides can used as a ‘clock’ for:
  Archeology/paleoclimate (14C 5.7 x 103 yr)
  Earth sciences/geomorphology – 10Be (1.38 x 106 yr), 26Al (717 x 103 yr)
  Environmental/hydrology — 36Cl (300 x 103 yr), 129I (15.7 x106 yr)
  Biomedicine — 14C, 41Ca (99 x 103 yr)
  Nuclear science/oceanography — actinides (Pu isotopes, U, etc.)
1

## Slide 4
AMS – Why do we do it?
  AMS is an ultra-sensitive method to detect trace amounts of long-lived radionuclides.
  With both cosmogenic and anthropogenic sources, these nuclides can used as a ‘clock’ for:
  Archeology/paleoclimate (14C 5.7 x 103 yr)
  Earth sciences/geomorphology – 10Be (1.38 x 106 yr), 26Al (717 x 103 yr)
  Environmental/hydrology — 36Cl (300 x 103 yr), 129I (15.7 x106 yr)
  Biomedicine — 14C, 41Ca (99 x 103 yr)
  Nuclear science/oceanography — actinides (Pu isotopes, U, etc.)
1
Decay counting counts atoms that decay over time: grams of sample, days of counting

## Slide 5
AMS – Why do we do it?
  AMS is an ultra-sensitive method to detect trace amounts of long-lived radionuclides.
  With both cosmogenic and anthropogenic sources, these nuclides can used as a ‘clock’ for:
  Archeology/paleoclimate (14C 5.7 x 103 yr)
  Earth sciences/geomorphology – 10Be (1.38 x 106 yr), 26Al (717 x 103 yr)
  Environmental/hydrology — 36Cl (300 x 103 yr), 129I (15.7 x106 yr)
  Biomedicine — 14C, 41Ca (99 x 103 yr)
  Nuclear science/oceanography — actinides (Pu isotopes, U, etc.)
Decay counting counts atoms that decay over time: grams of sample, days of counting
AMS utilizes direct detection – potentially every atom can be extracted and detected: mg of sample, minutes of counting
1

## Slide 6: The main challenge of (A)MS? The isobar
2

## Slide 7
AMS – how do we do it?
3

## Slide 8
AMS – how do we do it? (isobar suppression)
Eion
Suppressed negative ion formation
Molecular destruction
dE/dx separation
3

## Slide 9
AMS – how do we do it?
3

## Slide 10
AMS – how do we do it? (compact AMS)
  The working solution at ETH Zürich: compact AMS.
  Use of He as stripper gas allows lower energies  better transmission
  200kV MICADAS (46%) (12C1+– 50%)
  50kV LEA (50%) (12C1+ – 60%)
4

## Slide 11
AMS – how do we do it? (compact AMS)
2582
3195
5

## Slide 12
AMS – how do we do it? (compact AMS)
3195
Ion Source
(negative ion production)
LE Magnet
(mass selection)
200 kV Accelerator
(electron stripping, molecular destruction)
HE Magnet
(charge state selection)
HE ESA
(stigmatic correction)
5

## Slide 13
Good efficiency – both sputtering efficiency and extraction, so that we can use the smallest amount of sample possible and that measurements do not take too long
  1 mg of carbon
Good stability – so that optimum tuning parameters further downstream do not drift too far during the measurement
Ideally also: low maintenance!
What do we need from the source?
6

## Slide 14
What is a Cs sputter source?
K. Fifield, The Methodology and Physics of Accelerator Mass Spectrometry (2025)
7

## Slide 15
Vacuum lock sample ‘magazine’
Automatic sample changer
HV below
ETH source design

## Slide 16
Ø 127
Ø 191.3
Extractor lens
Ioniser
Target/cathode (on potential)
Cs nozzle
Permanent magnets
Isolating ceramics (Macor)
8

## Slide 17
Ø 127
Ø 191.3
Extractor lens
Ioniser
Target/cathode (on potential)
Cs nozzle
Permanent magnets
Isolating ceramics (Macor)
8

## Slide 18
Cu block for Cs heating
-Cs reservoir inside (5g Cs)
- air-cooled (1/4” pipe)
Ioniser platform (Cu wire for heating ioniser, generally operated at 15 A, 6V  90 W)
Target/cathode
(-6 kV )
Ceramic spacers between each voltage platform, with shields to protect from excessive carbon build up
Cs nozzle
Extraction potential (up to 10 kV)

## Slide 19
ETH source parameters and performance
‘Typical’ operation:
Ioniser = -36 kV
Target = -5.5 kV
Extractor = 5.6 kV
Corresponds to ~10s keV injection energy
“If the machine is running 24/7 (30uA 12C at the high energy end), then the source needs to be cleaned every 45-60 days”
9

## Slide 20
ETH source parameters and performance
7 oxa2 standards (only 3 shown), close to 2 M counts each in 8500 s (2.3 hrs) each
14C/12C x10-12
1.91 M
1.96 M
1.96 M
10

## Slide 21
Why push to even lower energy?
36S-
(2.08eV)
36Cl-
(3.61eV)
532nm
Nsurviving=N0⋅e(-σϕt)
σ: detachment cross section
ϕ: photon flux
t: interaction time
(It’s isobar suppression again)
11

## Slide 22
M. Martischini et al, IJMS 415, 9-17 (2017)
Why push to even lower energy?
36S-
(2.08eV)
36Cl-
(3.61eV)
532nm
11

## Slide 23
M. Martischini et al, IJMS 415, 9-17 (2017)
Why push to even lower energy?
Reason 1: ease of RFQ construction
36S-
(2.08eV)
36Cl-
(3.61eV)
532nm
11

## Slide 24
Why push to even lower energy?
2582
3195
For a maximum field of 0.695 T at LE magnet:
36Cl- = 40 keV
41CaH- = 32 keV
41CaF3- = 14 keV (!!!)
Reason 2: existing LE magnet limit
12

## Slide 25
Why push to even lower energy?
Reason 1: ease of RFQ construction
Reason 2: existing LE magnet limit
Reason 3: utilize the benefits of compact AMS for more isotopes
13

## Slide 26
Einzel Lens
RFQ
LE ESA
Diagnostics
Reacceleration
Deceleration
35Cl
37Cl
36Cl
14

## Slide 27
What parameters can we optimise?
For the same total energy, we can independently control:
Target/sputter potential and ioniser potential
Ioniser-sample distance
Cs temperature
Plus more (!)
15

## Slide 28
What parameters can we optimise?
For the same total energy, we can independently control:
Target/sputter potential and ioniser potential
Ioniser-sample distance
Cs temperature
Plus more (!)
15

## Slide 29: How were these measurements made?
Cs sputter ion source
LE magnet
BPM
Beam Dump
Einzel (box) Lens
12C-
16

## Slide 30: How were these measurements made?
Cs sputter ion source
LE magnet
BPM
Beam Dump
Einzel (box) Lens
12C-
16

## Slide 31
Target potential 4 kV
Ioniser potential 22 kV
Total energy 26 keV
5.7 mm
Closer to
ioniser
Further from
ioniser
-1.5mm
+1.5mm
First examination: target potential and distance
Examined using aluminium ‘dummy’ targets
                Optimum sputter spot (size, ‘sharpness’) found for fixed target potentials
17

## Slide 32
4kV sputter energy
2kV sputter energy
6kV sputter energy
Closer to
ioniser
Further from
ioniser
-1.5mm
+1.5mm
18

## Slide 33
4kV sputter energy
2kV sputter energy
6kV sputter energy
As sputter energy increases, the ‘ideal’ sputter spot is created when the sample is closer to the ioniser
Why?
Closer to
ioniser
Further from
ioniser
-1.5mm
+1.5mm
18

## Slide 34
4kV sputter energy
2kV sputter energy
6kV sputter energy
Sample + immersion lens
Ioniser + Cs nozzle
19

## Slide 35
4kV sputter energy
2kV sputter energy
6kV sputter energy
Sample + immersion lens
Ioniser + Cs nozzle
As sputter energy increases, gradient between ioniser and sample is stronger, bringing Cs+ to a focus in a shorter distance  sample needs to be moved closer
19

## Slide 36
120°C
130°C
110°C
More Cs temperature gives a greater Cs+ density, which ‘shields’ the target, hence needing a greater sputter energy to overcome it
Second examination: temperature and ionizer potential
20

## Slide 37
21
Notes: Should be able to see extractor current change CHECK WITH GRAFANA AT TIMES OF MEASUREMENT

## Slide 38
Seem to be extractor limited at the low energies
  That is, the lens formed between the extractor and the ioniser is not appropriate to extract all of the beam formed at these energies
21
Notes: Should be able to see extractor current change CHECK WITH GRAFANA AT TIMES OF MEASUREMENT

## Slide 39
Seem to be extractor limited at the low energies
  That is, the lens formed between the extractor and the ioniser is not appropriate to extract all of the beam formed at these energies
The other possibility is that the phase space is so large, that the beam is hitting the magnet chamber
21
Notes: Should be able to see extractor current change CHECK WITH GRAFANA AT TIMES OF MEASUREMENT

## Slide 40
22

## Slide 41
22

## Slide 42
Cs sputter ion source
LE magnet
Object aperture
Image aperture
BPM
Beam Dump
Einzel (box) Lens
What about the phase space?
23

## Slide 43
Cs sputter ion source
LE magnet
Object aperture
Image aperture
BPM
Beam Dump
Einzel (box) Lens
What about the phase space?
* Calculations assume image aperture is size-defining
23

## Slide 44
What about the phase space?
24

## Slide 45
What about the phase space?
Could we be sampling only part of the phase space?
24

## Slide 46
What about the phase space?
Could we be sampling only part of the phase space?
24

## Slide 47
What did we learn? - Conclusions
Low energy operation seems to be extraction-optics limited
Beam quality doesn’t behave as expected – phase space vs magnet chamber?
25

## Slide 48
Diagnostics and measurement technique:
  Quad-scan emittance using box lens
  Change aperture + BPM setup
Ion optics:
  Introduce negative extractor potential
Open physics question:
  Is the low-energy limitation dominated by re-neutralization or phase-space blow-up?
What next? – Future work
26

## Slide 49
Diagnostics and measurement technique:
  Quad-scan emittance using box lens
  Change aperture + BPM setup
Ion optics:
  Introduce negative extractor potential
Open physics question:
  Is the low-energy limitation dominated by re-neutralization or phase-space blow-up?
What next? – Future work
26

## Slide 50
Acknowledgements
Laboratory of Ion Beam Physics technical staff, and the many semester students and apprentices
Thank you for you attention!
27

## Slide 51: ETH Cs sputter source

## Slide 52
Oort
Wolf
Spörer
Maunder
Dalton

## Slide 53: ETH source performance - temperature
Chart: Target voltage [kV]
| 1 | 1.5 | 2 | 2.5 | 3 | 3.5 | 4 | 4.5 | 5 | 5.5 | 6 | 6.5 | 7 | 7.5 | 8 | 8.5 | 9 | 9.5 | 10 | 10.5
100 degC | 3.8 | 12.8 | 18.6 | 26.2 | 27.2 | 25 | 23.1 | 20.9 | 19.6 | 18.7 | 17.9 | 17.6 | 17.4 | 17.3 | 17.3 | 17.2 | 17 | 16.9 | 16.5 | 16.6
105 degC | 4.1 | 13.4 | 21 | 28.1 | 36.6 | 42.2 | 40.3 | 36.6 | 34 | 31.6 | 29.5 | 28.2 | 27.3 | 26.4 | 25.7 | 25.1 | 24.6 | 24.2 | 23.9 | 23.6
110 deg C | 4.1 | 14.1 | 24.8 | 34.6 | 42 | 51.2 | 62.3 | 62.1 | 58.2 | 54.1 | 49.9 | 47 | 44.6 | 42.4 | 40.6 | 38.9 | 37.5 | 36 | 34.6 | 33.2
115 degC | 4.6 | 15 | 26.2 | 38.7 | 51.3 | 64 | 73.8 | 86.5 | 97.8 | 98.1 | 93.1 | 87.4 | 81.6 | 76.5 | 72.7 | 68.8 | 65.6 | 63.5 | 60.6 | 58.2
120 degC | 4.4 | 14 | 24.7 | 38.1 | 53.4 | 71.2 | 83.4 | 96.8 | 107.2 | 124.1 | 135.8 | 140.5 | 138.2 | 130.7 | 125.8 | 118.8 | 114.2 | 110.3 | 104.7 | 101.4
125 degC | 4.7 | 15.1 | 26.6 | 40.2 | 55.8 | 77.3 | 92.3 | 109.3 | 122 | 134.2 | 146.8 | 158.7 | 179.2 | 186.6 | 190.5 | 187.5 | 180.5 | 170.6 | 160.9 | 152.1
130 degC | 4.8 | 14.7 | 27.1 | 41.3 | 56.8 | 80 | 99.4 | 124 | 141.6 | 157.5 | 175.9 | 191.9 | 213 | 226 | 245 | 265 | 272 | 274 | 276.3 | 271

## Slide 54: Compact AMS - Cl charge states
Charge state yield for Cl+ around 50% for compact accelerator energies (150-200 keV), while high energy tandem energies (46 MeV) requires use of a higher charge state 7+, which gives a much lower yield, around 15%
23
Measurements by L. Danini, semester project

## Slide 55
RFQ + LPD: magnet limited for the higher masses (example for 36Cl, 41Ca), construction of RFQ, optimum ratio of accelerator to source potential (where we might want to operate - ~100kV accelerator)
  200kV MICADAS (46%) (CSY – 50%)
  50kV LEA (50%) (CSY – 60%)
    Basically, better performance as you go down in energy (transmission wins in particular) but compromise on molecular destruction and angular spread/scatter
Big advantage to move to low energy for Cl
He advantage – anomaly in the 1+ CSY due to lack of ‘back reaction’/neutralisation reaction
Ultimate question: how low can we operate the source while retaining a useful, transportable beam?
  i.e. small/reasonable phase space, good transmission
For a maximum field of 0.695 T at LE magnet
36Cl- = 40 keV
41CaH- = 32 keV
41CaF3- = 14 keV (!!!)
M. Martischini et al, IJMS 415, 9-17 (2017)
Why push to even lower energy?